Fix **Low Water Pressure** Inhibit

Inadvertently pressing the red button on the side of the couch controls can cause the Low Water Pressure inhibit, which prevents machine operation (running beam, moving the couch, pretty much everything). To remove the inhibit:

1. Press the red button on the side of either couch control.
2. Press the reset motors button from the console outside the treatment room.
3. Press the **RESET MOTORS** button on the hand pendant inside the treatment room.